Supplementary material

Table S1: Stock solution for the synthetic standard nutrient mixture (1000mg/L of targeted ions in DI water).

| Sr. no. | Compound | Concentration | Target ion |
| --- | --- | --- | --- |
| 1 | Sodium phosphate dibasic (Na₂HPO₄) | 12.50 g/L | HPO₄²⁻ (1000mg/L) |
| 2 | Sodium nitrate (NaNO₃) | 60.66 g/L | NO₃⁻-N (1000mg/L) |
| 3 | Ammonium chloride (NH₄Cl) | 38.18 g/L | NH₄⁺-N (1000mg/L) |
| 4 | Calcium chloride dihydrate (CaCl₂⋅2H₂O) | 36.68 g/L | Ca²⁺ (1000mg/L) |
| 5 | Potassium chloride (KCl) | 19.07 g/L | K+ (1000mg/L) |

Table S2: Synthetic standard nutrient mixture (test samples) for interference testing.

| Sample ID | Target HPO42−​ (mg/L) | Target Ca2+ (mg/L) | Target K+ (mg/L) | Target NH4+​-N (mg/L) | Target NO3−​-N (mg/L) |
| --- | --- | --- | --- | --- | --- |
| Mix-01 | 16.50 | 300.00 | 400.00 | 2.00 | 290.00 |
| Mix-02 | 330.00 | 15.00 | 20.00 | 30.00 | 140.00 |
| Mix-03 | 50.00 | 280.00 | 180.00 | 10.00 | 250.00 |
| Mix-04 | 180.00 | 40.00 | 310.00 | 25.00 | 160.00 |
| Mix-05 | 90.00 | 210.00 | 60.00 | 8.00 | 220.00 |
| Mix-06 | 250.00 | 100.00 | 220.00 | 12.00 | 190.00 |

Table S3: Low and high calibration standard solutions for synthetic standard nutrient mixture samples testing.

| Ion | Low [mg/L] | High [mg/L] |
| --- | --- | --- |
| HPO42- | 16.50 | 330.00 |
| NO3- | 88.00 | 1771.00 |
| Ca2+ | 15.00 | 300.00 |
| NH4+ | 25.00 | 514.00 |
| K+ | 20.00 | 400.00 |

Table S4: Chemical composition of the complex commercial nutrient mixture

| Nutrient Element | Calculated value (Low calibration solution-0.5mS/cm) | Calculated value (High calibration solution 2.5mS/cm) |
| --- | --- | --- |
|  | mg/L | mg/L |
| N (NO3) | 44.49 | 222.49 |
| N(NH4+) | 4.23 | 21.17 |
| HPO42- | 18.85 | 94.25 |
| K+ | 78.90 | 394.50 |
| Ca2+ | 36.45 | 182.25 |

Formulas:

- Relative Standard Deviation

%RSD = (Standard deviation/Mean) *100

Where:

Standard Deviation): The scatter or spread of your individual data points around the average.

Mean: The average value of all your repeated measurements.

- Percent Recovery (%)

% R = (X measured / X expected) *100

Where:

X Measured: It is the measured value of ions obtained through ICP-MS, colorimetric kits, or AutoNutri methods.

X Expected: It is the expected value of ions obtained through calculations.

- Relative error (%)

%RE = (X measured – X expected /X expected) * 100

Where:

X Measured: It is the measured value of ions obtained through ICP-MS, colorimetric kits, or AutoNutri methods.

X Expected: It is the expected value of ions obtained through calculations.

- Testing & Validation for Standard Nutrient Mixture

Table S5: Comparison of the ICP-MS and colorimetric kit-based measured values with the expected concentration values for each standard ion in the synthetic standard nutrient mixture (mean± standard deviation).

| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-01_R1 | MIX-01_R2 | MIX-01_R3 | Average (mg/L) | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Colorimetric method | NH₄⁺-N | 2.00 | 5.00 | 5.00 | 7.00 | 5.67 | 1.15 | 20.38 | 283.33 | 183.33 |
|  | NO₃⁻-N | 290.00 | 250.00 | 300.00 | 350.00 | 300.00 | 50.00 | 16.67 | 103.45 | 3.45 |
| ICP-MS method | HPO₄²⁻ | 16.50 | 18.09 | 18.40 | 17.66 | 18.05 | 0.37 | 2.05 | 109.38 | 9.38 |
|  | K+ | 400.00 | 445.00 | 464.00 | 457.00 | 455.33 | 9.61 | 2.11 | 113.83 | 13.83 |
|  | Ca2+ | 300.00 | 312.00 | 321.00 | 315.00 | 316.00 | 4.58 | 1.45 | 105.33 | 5.33 |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-02_R1 | MIX-02_R2 | MIX-02_R3 | Average (mg/L) | Std deviation | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| Colorimetric method | NH₄⁺-N | 30.00 | 25.00 | 40.00 | 30.00 | 31.67 | 7.64 | 24.12 | 105.56 | 5.56 |
|  | NO₃⁻-N | 140.00 | 150.00 | 150.00 | 170.00 | 156.67 | 11.55 | 7.37 | 111.90 | 11.90 |
| ICP-MS method | HPO₄²⁻ | 330.00 | 364.85 | 358.72 | 370.99 | 364.85 | 6.13 | 1.68 | 110.56 | 10.56 |
|  | K+ | 20.00 | 21.30 | 21.30 | 21.50 | 21.37 | 0.12 | 0.54 | 106.83 | 6.83 |
|  | Ca2+ | 15.00 | 16.00 | 15.80 | 17.40 | 16.40 | 0.87 | 5.32 | 109.33 | 9.33 |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-03_R1 | MIX-03_R2 | MIX-03_R3 | Average (mg/L) | Std deviation | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| Colorimetric method | NH₄⁺-N | 10.00 | 10.00 | 13.00 | 20.00 | 14.33 | 5.13 | 35.80 | 143.33 | 43.33 |
|  | NO₃⁻-N | 250.00 | 250.00 | 250.00 | 300.00 | 266.67 | 28.87 | 10.83 | 106.67 | 6.67 |
| ICP-MS method | HPO₄²⁻ | 50.00 | 53.04 | 53.04 | 52.43 | 52.84 | 0.35 | 0.67 | 105.67 | 5.67 |
|  | K+ | 180.00 | 185.00 | 187.00 | 189.00 | 187.00 | 2.00 | 1.07 | 103.89 | 3.89 |
|  | Ca2+ | 280.00 | 288.00 | 290.00 | 293.00 | 290.33 | 2.52 | 0.87 | 103.69 | 3.69 |

| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-04_R1 | MIX-04_R2 | MIX-04_R3 | Average (mg/L) | Std deviation | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Colorimetric method | NH₄⁺-N | 25.00 | 25.00 | 25.00 | 30.00 | 26.67 | 2.89 | 10.83 | 106.67 | 6.67 |
|  | NO₃⁻-N | 160.00 | 150.00 | 150.00 | 190.00 | 163.33 | 23.09 | 14.14 | 102.08 | 2.08 |
| ICP-MS method | HPO₄²⁻ | 180.00 | 189.48 | 186.11 | 181.20 | 185.59 | 4.16 | 2.24 | 103.11 | 3.11 |
|  | K+ | 310.00 | 345.00 | 351.00 | 344.00 | 346.67 | 3.79 | 1.09 | 111.83 | 11.83 |
|  | Ca2+ | 40.00 | 41.80 | 43.20 | 41.30 | 42.10 | 0.98 | 2.34 | 105.25 | 5.25 |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-05_R1 | MIX-05_R2 | MIX-05_R3 | Average (mg/L) | Std deviation | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| Colorimetric method | NH₄⁺-N | 8.00 | 10.00 | 10.00 | 13.00 | 11.00 | 1.73 | 15.75 | 137.50 | 37.50 |
|  | NO₃⁻-N | 220.00 | 200.00 | 200.00 | 250.00 | 216.67 | 28.87 | 13.32 | 98.48 | -1.52 |
| ICP-MS method | HPO₄²⁻ | 90.00 | 93.51 | 94.13 | 91.98 | 93.21 | 1.11 | 1.19 | 103.56 | 3.56 |
|  | K+ | 60.00 | 62.20 | 62.40 | 62.90 | 62.50 | 0.36 | 0.58 | 104.17 | 4.17 |
|  | Ca2+ | 210.00 | 210.00 | 217.00 | 220.00 | 215.67 | 5.13 | 2.38 | 102.70 | 2.70 |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-06_R1 | MIX-06_R2 | MIX-06_R3 | Average (mg/L) | Std deviation | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| Colorimetric method | NH₄⁺-N | 12.00 | 10.00 | 10.00 | 15.00 | 11.67 | 2.89 | 24.74 | 97.22 | -2.78 |
|  | NO₃⁻-N | 190.00 | 180.00 | 200.00 | 200.00 | 193.33 | 11.55 | 5.97 | 101.75 | 1.75 |
| ICP-MS method | HPO₄²⁻ | 250.00 | 281.15 | 278.70 | 296.48 | 285.44 | 9.64 | 3.38 | 114.18 | 14.18 |
|  | K+ | 220.00 | 243.00 | 245.00 | 254.00 | 247.33 | 5.86 | 2.37 | 112.42 | 12.42 |
|  | Ca2+ | 100.00 | 88.50 | 90.30 | 93.80 | 90.87 | 2.70 | 2.97 | 90.87 | -9.13 |

Table S6: Comparison of the AutoNutri measured values with the expected concentration values for each standard ion in the synthetic standard nutrient mixture (mean± standard deviation).

| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-01_R1 | MIX-01_R2 | MIX-01_R3 | Average | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| AutoNutri method | NH₄⁺-N | 2.00 | 2.31 | 2.74 | 2.65 | 2.57 | 0.22 | 8.67 | 128.50 | 28.29 |
|  | NO₃⁻-N | 290.00 | 278.33 | 291.72 | 289.63 | 286.56 | 7.21 | 2.52 | 98.81 | -1.19 |
|  | HPO₄²⁻ | 16.50 | 14.34 | 22.44 | 26.75 | 21.18 | 6.30 | 29.74 | 128.36 | 28.35 |
|  | K+ | 400.00 | 384.96 | 397.06 | 425.81 | 402.61 | 20.98 | 5.21 | 100.65 | 0.65 |
|  | Ca2+ | 300.00 | 314.01 | 286.77 | 308.57 | 303.12 | 14.41 | 4.75 | 101.04 | 1.04 |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-02_R1 | MIX-02_R2 | MIX-02_R3 | Average | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| AutoNutri method | NH₄⁺-N | 30.00 | 27.28 | 28.39 | 27.60 | 27.75 | 0.57 | 2.06 | 92.50 | -7.48 |
|  | NO₃⁻-N | 140.00 | 155.42 | 164.53 | 147.32 | 155.76 | 8.61 | 5.53 | 111.26 | 11.25 |
|  | HPO₄²⁻ | 330.00 | 310.39 | 370.33 | 363.84 | 348.19 | 32.89 | 9.45 | 105.51 | 5.51 |
|  | K+ | 20.00 | 29.39 | 31.17 | 31.72 | 30.76 | 1.22 | 3.95 | 153.80 | 53.79 |
|  | Ca2+ | 15.00 | 24.06 | 29.82 | 23.66 | 25.85 | 3.45 | 13.35 | 172.33 | 72.31 |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-03_R1 | MIX-03_R2 | MIX-03_R3 | Average | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| AutoNutri method | NH₄⁺-N | 10.00 | 9.77 | 11.57 | 7.34 | 9.56 | 2.13 | 22.25 | 95.60 | -4.40 |
|  | NO₃⁻-N | 250.00 | 233.29 | 271.82 | 269.76 | 258.29 | 21.68 | 8.39 | 103.32 | 3.32 |
|  | HPO₄²⁻ | 50.00 | 49.51 | 54.05 | 57.13 | 53.56 | 3.84 | 7.16 | 107.12 | 7.12 |
|  | K+ | 180.00 | 136.74 | 166.84 | 189.84 | 164.47 | 26.63 | 16.19 | 91.37 | -8.63 |
|  | Ca2+ | 280.00 | 266.37 | 272.43 | 283.81 | 274.20 | 8.85 | 3.23 | 97.93 | -2.07 |

| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-04_R1 | MIX-04_R2 | MIX-04_R3 | Average | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| AutoNutri method | NH₄⁺-N | 25.00 | 19.79 | 27.49 | 22.48 | 23.25 | 3.91 | 16.81 | 93.00 | -6.99 |
|  | NO₃⁻-N | 160.00 | 173.91 | 161.76 | 182.66 | 172.78 | 10.49 | 6.07 | 107.99 | 7.99 |
|  | HPO₄²⁻ | 180.00 | 189.35 | 167.05 | 167.83 | 174.74 | 12.66 | 7.24 | 97.08 | -2.92 |
|  | K+ | 310.00 | 298.18 | 282.04 | 287.90 | 289.38 | 8.17 | 2.82 | 93.35 | -6.65 |
|  | Ca2+ | 40.00 | 40.39 | 43.99 | 58.71 | 47.70 | 9.70 | 20.34 | 119.25 | 19.24 |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-05_R1 | MIX-05_R2 | MIX-05_R3 | Average | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| AutoNutri method | NH₄⁺-N | 8.00 | 10.56 | 9.00 | 10.71 | 10.09 | 0.95 | 9.37 | 126.13 | 26.15 |
|  | NO₃⁻-N | 220.00 | 226.45 | 220.42 | 200.42 | 215.76 | 13.62 | 6.31 | 98.07 | -1.93 |
|  | HPO₄²⁻ | 90.00 | 70.26 | 106.68 | 82.28 | 86.41 | 18.56 | 21.48 | 96.01 | -3.99 |
|  | K+ | 60.00 | 60.71 | 66.24 | 60.45 | 62.47 | 3.27 | 5.24 | 104.12 | 4.11 |
|  | Ca2+ | 210.00 | 191.05 | 189.16 | 197.74 | 192.65 | 4.50 | 2.34 | 91.74 | -8.26 |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-06_R1 | MIX-06_R2 | MIX-06_R3 | Average | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| AutoNutri method | NH₄⁺-N | 12.00 | 11.18 | 13.86 | 16.28 | 13.77 | 2.55 | 18.53 | 114.75 | 14.76 |
|  | NO₃⁻-N | 190.00 | 206.61 | 194.90 | 197.52 | 199.68 | 6.15 | 3.08 | 105.09 | 5.09 |
|  | HPO₄²⁻ | 250.00 | 241.17 | 234.00 | 237.62 | 237.60 | 3.59 | 1.51 | 95.04 | -4.96 |
|  | K+ | 220.00 | 229.56 | 208.78 | 232.18 | 223.51 | 12.82 | 5.74 | 101.60 | 1.59 |
|  | Ca2+ | 100.00 | 115.79 | 105.68 | 95.46 | 105.64 | 10.16 | 9.62 | 105.46 | 5.64 |

- Testing & Validation for Commercial Nutrient Mixture

Table S7: Comparison of the ICP-MS and colorimetric kit-based measured values with the expected concentration values for each standard ion in the complex commercial nutrient mixture.

| Complex commercial nutrient mixture-A (EC 1.0 mS/cm) |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-A_R1 (mg/L) | MIX-A_R2 (mg/L) | MIX-A_R3 (mg/L) | Average (mg/L) | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) |
| Colorimetric method | NH₄⁺-N | 8.47 | 25.00 | 13.00 | 10.00 | 16.00 | 7.94 | 49.61 | 188.90 | 88.90 |
|  | NO₃⁻-N | 88.99 | 100.00 | 50.00 | 100.00 | 83.33 | 28.87 | 34.64 | 93.64 | -6.36 |
| ICP-MS method | HPO₄²⁻ | 37.70 | 31.60 | 40.10 | 56.90 | 42.87 | 12.87 | 30.03 | 113.70 | 13.70 |
|  | K+ | 157.80 | 128.00 | 162.00 | 229.00 | 173.00 | 51.39 | 29.71 | 109.63 | 9.63 |
|  | Ca2+ | 72.90 | 69.30 | 88.20 | 125.00 | 94.17 | 28.33 | 30.08 | 129.17 | 29.17 |
| Complex commercial nutrient mixture-B (EC 1.4 mS/cm) |  |  |  |  |  |  |  |  |  |  |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-B_R1 (mg/L) | MIX-B_R2 (mg/L) | MIX-B_R3 (mg/L) | Average (mg/L) | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) |
| Colorimetric method | NH₄⁺-N | 11.85 | 30.00 | 10.00 | 10.00 | 16.67 | 11.55 | 69.28 | 140.65 | 40.65 |
|  | NO₃⁻-N | 124.58 | 100.00 | 100.00 | 150.00 | 116.67 | 28.87 | 24.74 | 93.65 | -6.35 |
| ICP-MS method | HPO₄²⁻ | 52.80 | 32.70 | 45.80 | 61.20 | 46.57 | 14.27 | 30.63 | 88.19 | -11.81 |
|  | K+ | 220.90 | 129.00 | 177.00 | 245.00 | 183.67 | 58.29 | 31.74 | 83.14 | -16.86 |
|  | Ca2+ | 102.00 | 69.80 | 101.00 | 135.00 | 101.93 | 32.61 | 31.99 | 99.93 | -0.07 |
| Complex commercial nutrient mixture-C (EC 1.8 mS/cm) |  |  |  |  |  |  |  |  |  |  |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-C_R1 (mg/L) | MIX-C_R2 (mg/L) | MIX-C_R3 (mg/L) | Average (mg/L) | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) |
| Colorimetric method | NH₄⁺-N | 15.24 | 25.00 | 13.00 | 10.00 | 16.00 | 7.94 | 49.61 | 104.99 | 4.99 |
|  | NO₃⁻-N | 160.18 | 150.00 | 150.00 | 160.00 | 153.33 | 5.77 | 3.77 | 95.73 | -4.27 |
| ICP-MS method | HPO₄²⁻ | 67.90 | 40.50 | 53.60 | 60.80 | 51.63 | 10.29 | 19.93 | 76.04 | -23.96 |
|  | K+ | 284.00 | 160.00 | 220.00 | 245.00 | 208.33 | 43.68 | 20.97 | 73.36 | -26.64 |
|  | Ca2+ | 131.20 | 87.60 | 117.00 | 134.00 | 112.87 | 23.47 | 20.80 | 86.03 | -13.97 |

Table S8: Comparison of the AutoNutri system-measured values with the expected concentration values for each standard ion in the complex commercial nutrient mixture.

| Complex commercial nutrient mixture-A (EC 1.0 mS/cm) |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-A_R1 (mg/L) | MIX-A_R2 (mg/L) | MIX-A_R3 (mg/L) | Average (mg/L) | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (AutoNutri against expected values) |
| AutoNutri | NH₄⁺-N | 8.47 | 8.77 | 9.13 | 10.31 | 9.40 | 0.81 | 8.57 | 111.02 | 11.02 |
| method | NO₃⁻-N | 88.99 | 89.41 | 93.74 | 86.87 | 90.01 | 3.47 | 3.86 | 101.14 | 1.14 |
|  | HPO₄²⁻ | 37.7 | 42.08 | 39.06 | 36.46 | 39.20 | 2.81 | 7.18 | 103.98 | 3.98 |
|  | K+ | 157.8 | 179.45 | 177.17 | 174.02 | 176.88 | 2.73 | 1.54 | 112.09 | 12.09 |
|  | Ca2+ | 72.9 | 83.75 | 76.68 | 74.88 | 78.44 | 4.69 | 5.98 | 107.59 | 7.59 |
| Complex commercial nutrient mixture-B (EC 1.4 mS/cm) |  |  |  |  |  |  |  |  |  |  |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-B_R1 (mg/L) | MIX-B_R2 (mg/L) | MIX-B_R3 (mg/L) | Average (mg/L) | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| AutoNutri | NH₄⁺-N | 11.85 | 15.73 | 15.89 | 14.56 | 15.39 | 0.72 | 4.71 | 129.91 | 29.91 |
| method | NO₃⁻-N | 124.58 | 143.36 | 121.77 | 139.89 | 135.01 | 11.59 | 8.59 | 108.37 | 8.37 |
|  | HPO₄²⁻ | 52.8 | 68.74 | 53.56 | 58.54 | 60.28 | 7.74 | 12.84 | 114.16 | 14.16 |
|  | K+ | 220.9 | 221.37 | 224.42 | 219.72 | 221.84 | 2.39 | 1.08 | 100.42 | 0.42 |
|  | Ca2+ | 102 | 113.01 | 110.36 | 102.54 | 108.64 | 5.44 | 5.01 | 106.51 | 6.51 |
| Complex commercial nutrient mixture-C (EC 1.8 mS/cm) |  |  |  |  |  |  |  |  |  |  |
| Analysis method | Element Present | Expected Conc. of target element (mg/L) | MIX-C_R1 (mg/L) | MIX-C_R2 (mg/L) | MIX-C_R3 (mg/L) | Average (mg/L) | Std deviation (SD) | Relative standard deviation (%) | Percent Recovery (%) | Relative error (%) (measured values Vs expected values) |
| AutoNutri | NH₄⁺-N | 15.24 | 14.84 | 17.49 | 16.75 | 16.36 | 1.37 | 8.38 | 107.35 | 7.35 |
| method | NO₃⁻-N | 160.18 | 141.98 | 140.73 | 154.53 | 145.74 | 7.63 | 5.24 | 90.99 | -9.01 |
|  | HPO₄²⁻ | 67.9 | 81.04 | 74.32 | 77.84 | 77.73 | 3.36 | 4.32 | 114.48 | 14.48 |
|  | K+ | 284 | 295.04 | 297.82 | 281.29 | 291.38 | 8.85 | 3.04 | 102.60 | 2.60 |
|  | Ca2+ | 131.2 | 137.27 | 121.56 | 149.36 | 136.06 | 13.94 | 10.25 | 103.71 | 3.71 |

Table S9: Comparison of the AutoNutri measured values with reference method for each standard ion in the complex commercial nutrient mixture (Mean± Standard deviation).

| Element Present | Total Expected Conc. of target element (mg/L) | ICP-MS measured value (mg/L) | Colorimetric kit-based method measured value (mg/L) | AutoNutri measured value (mg/L) | Recovery of Autonutri (%) against reference method | Absolute error (AutoNutri Vs reference method) | Relative error % (AutoNutri Vs reference method) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Complex commercial nutrient Mixture A (EC 1.0mS/cm) |  |  |  |  |  |  |  |
| NH₄⁺-N | 8.47 | N/A | 16.00±7.94 | 9.40±0.81 | 58.75 | -6.60 | -41.25 |
| NO₃⁻-N | 88.99 | N/A | 83.33±28.87 | 90.01±3.47 | 108.02 | 6.67 | 8.00 |
| HPO₄²⁻ | 37.70 | 42.87±12.87 | N/A | 39.20±2.81 | 91.44 | -3.67 | -8.56 |
| K+ | 157.80 | 173.00±51.39 | N/A | 176.88±2.73 | 102.24 | 3.88 | 2.24 |
| Ca2+ | 72.90 | 94.17±28.33 | N/A | 78.44±4.69 | 83.30 | -15.74 | -16.71 |
| Complex commercial nutrient Mixture B (EC 1.4mS/cm) |  |  |  |  |  |  |  |
| NH₄⁺-N | 11.85 | N/A | 16.67±11.55 | 15.39±0.72 | 92.32 | -1.28 | -7.68 |
| NO₃⁻-N | 124.58 | N/A | 116.67±28.87 | 135.01±11.59 | 115.72 | 18.34 | 15.72 |
| HPO₄²⁻ | 52.80 | 46.57±14.27 | N/A | 60.28±7.74 | 129.42 | 13.70 | 29.42 |
| K+ | 220.90 | 183.67±58.29 | N/A | 221.84±2.39 | 120.78 | 38.16 | 20.78 |
| Ca2+ | 102.00 | 101.93±32.61 | N/A | 108.64±5.44 | 106.58 | 6.70 | 6.57 |
| Complex commercial nutrient Mixture C (EC 1.8mS/cm) |  |  |  |  |  |  |  |
| NH₄⁺-N | 15.24 | N/A | 16.00±7.94 | 16.36±1.37 | 102.25 | 0.36 | 2.25 |
| NO₃⁻-N | 160.18 | N/A | 153.33±5.77 | 145.74±7.63 | 95.05 | -7.59 | -4.95 |
| HPO₄²⁻ | 67.90 | 51.63±10.29 | N/A | 77.73±3.36 | 150.55 | 26.10 | 50.55 |
| K+ | 284.00 | 208.33±43.68 | N/A | 291.38±8.85 | 139.86 | 83.05 | 39.86 |
| Ca2+ | 131.20 | 112.87±23.47 | N/A | 136.06±13.94 | 120.55 | 23.19 | 20.55 |